[image: image1.jpg]-
Y
Stago

Diagnostica Stago, Inc.
Five Century Drive
Parsippany, N) 07054
USA

1-800-222-COAG
973-631-1200

Fax: 973-631-1618
www.stago-us.com

STAGO - TDL US 2159x2794.indd 1 4/06/07 14:57:46





August 30, 2010
Re: Addendum to STA® Compact, STA® Compact CT, STA-R® and STA-R® Evolution

      Using Soft Stylets: Preventative Maintenance Optimization and Operator’s Security

Dear Valued Customer:

Please find attached an addendum that details a change in maintenance for your instrument. Stago continues to evaluate and improve our processes and has found that the continued use of metal stylets to clean out the needles on a weekly basis is not required.

As per the addendum, only use the soft stylet (Item number 80343) if you suspect a clog. Do not pass the stylet through a cap-piercing needle.

Sincerely,

Alicia Carpenter

Technical Training Manager

1-800-222-2624 x4198

MKT-10-022, Eff. Date 08/30/2010

